## Slide 1
High-accuracy Time and Frequency in VLBI
Katie Pazamickas
katherine.Pazamickas@exelisinc.com
Tom Clark
K3IO@verizon.net
Rick Hambly
Rick@cnssys.com

## Slide 2
Presentation Format
Background – Tom Clark
Oscillators and Clocks
What “Clock” Performance Does VLBI Need?
“Absolute Time” (i.e. Clock Accuracy )
The Hydrogen Maser - Katie Pazamickas
Maser Outputs
Data/Frequency Monitoring
Troubleshooting/Routine Maintenance
GPS Time - Rick Hambly
Week rollover may mean retiring old GPS receivers
GPS receiver’s quantization error
“Absolute” Receiver Calibration
New developments

## Slide 3
Long-Term
seconds - years
Events that occur with a defined
nsec -- minutes
Oscillator
Escapement Wheels & Pendulums
Crystal Oscillators
Cavity Oscillators
Oscillator Locked to Atomic Transition
  Rubidium (6.8 GHz)
  Cesium (9.1 GHz)
  Hydrogen Maser (1.4 GHz)
Integrator and Display = Clock
Gears
Electronic Counters
Real Clocks
FREQUENCY
TIMING
The Difference Between Frequency and Time
Oscillators and Clocks

## Slide 4: What “Clock” Performance Does VLBI Need?
The VLBI community (Radio Astronomy and Geodesy) uses Hydrogen Masers at 40-50 remote sites all around the world.
To achieve ~10° signal coherence for ~1000 seconds at 10 GHz we need the 2 clocks (oscillators) at the ends of the interferometer to maintain relative stability of:
   [10°/(360° * 1010Hz * 103sec)]
   2.8 * 10-15 @ 1000 sec.
1

## Slide 5
4
In Geodetic applications, the station clocks are modeled at relative levels ~30 psec over a day:
   [3010-12 / 86400 sec]
   3.510-16 @ 1 day
2
What “Clock” Performance Does VLBI Need?
Notes: 4

## Slide 6
5
To correlate data acquired at 16Mb/s, station timing at relative levels ~50 nsec or better is needed. After a few days of inactivity, this requires:
   [50 * 10-9 / 106 sec]
   5 * 10-14 @ 106 sec
Since VLBI now defines UT1, VLBI needs to control [UTC(USNO) - UTC(VLBI)] with an ACCURACY (traceable to USNO)
   100 nsec - 1 sec
To detect problems, VLBI should monitor the long-term behavior of the Hydrogen Masers (at least) every hour with PRECISION
  10-50 nsec
3
What “Clock” Performance Does VLBI Need?
Notes: 5

## Slide 7
FREQUENCY
TIME
1
2
3
Allan Deviations of Typical Clocks
Allan Deviation – A graphical look at clock performance
Katie Pazamickas, Tom Clark & Rick Hambly
@Haystack, May 2015

## Slide 8
Why do we need to worry about “Absolute Time” (i.e. Clock Accuracy ) in VLBI?
The ONLY real reason for worrying about “absolute time” is to relate the position of the earth to the position of the stars:
  Generating Sidereal Time to point antennas.
  Measuring UT1 (i.e. “Sundial Time”) to see changes due to redistribution of mass in/on the earth over long periods of time (a.k.a. “The Reference Frame”)
  Knowing the position of the earth with respect to the moon, planets and satellites.
  Making the correlation and Data Analysis jobs easier

## Slide 9
Why do we need to worry about “Absolute Time” (i.e. Clock Accuracy) in VLBI?
At the stations this means that we will need to pay more attention to timing elements like
  Frequency Standard and Station Timing
  The lengths of all signal & clock cables
  The geometry of the feed/receiver to the antenna.
  Calibration of instrumental delays inside the receiver and backend. The development of new instrumentation is needed.
  The care with which system changes are reported to the correlators and the data analysts.

## Slide 10
The Real Signal Path
VLBI Data Analysis assumes the Geometric Clock is at the Intersection of Axes of the Antenna *
* Note -- If the axes don’t intersect, then an “offset axis” model of the antenna is used
VLBI’s “REAL” Clocks (#1)

## Slide 11
H-Maser
Phase Cal Ground Unit: Monitors Cable Length Changes
Counter
Cable Length Transponder
Divide by n
5/10 MHz
Microwave Receiver
1/5/10 MHz
1 Pulse/μsec
DOWN
UP
This is the “clock” that is used to analyze VLBI data
Quasar
Pulse Generator
ON ANTENNA
CONTROL ROOM
IF
VLBI’s “REAL” Clocks (#2)

## Slide 12
IF From Microwave Receiver
IF Distributor
Down Converter
5/10 MHz
Clock in Mk5 or Mk6 (XCube) Formatter
5/10 MHz
Clipper/ Sampler
This is the “clock” the correlator uses to make fringes
Mark 5 or Mark 6 (XCube) Recorder
H-Maser
VLBI’s “REAL” Clocks (#3)

## Slide 13: Setting VLBI Clock Time & Rate with GPS
Compare two distant clocks by observing the same GPS satellite(s) at the same time (also called Common View)
  Requires some inter-visibility between sites
  Requires some near-real-time communication
  Links you directly to the “Master Clock” on the other end at ~1 nsec level
Use Geodetic GPS receivers (i.e. as an extension of the IGS network)
  Requires high quality, probably dual frequency, receiver but it’s hard to gain access to the internal clock.
  Requires transferring ~1 MB/day of data from site
  Requires fairly extensive computations using dual-frequency data to get ~300 psec results with ionosphere corrections
  Allows Geodetic community to use VLBI Site (and H-Maser) for geodesy
  Difficult to obtain “Real Time” clock pulses!
Blindly use the Broadcast GPS Timing Signals as a clock
  Yields “Real Time” ~10-30 nsec results with ~ $1000 hardware
  Single Frequency L1 only (for now) causes ionospheric error

## Slide 14: How we get less than 5 nsec 1-sigma timing
Start with a good timing receiver, like the Motorola Oncore or the Synergy SSR (uBlox).
Average the positioning data for ~1-2 days to determine the station’s coordinates. This should be good to <5 meters. Or if the site has been accurately surveyed, use the survey values.
Lock the receiver’s position to this average.
Make sure that your Time-Interval Counter (TIC) is triggering cleanly. Start the counter with the 1 PPS signal from the “house” atomic clock and stop with the GPS receiver’s 1PPS.
Average the individual one/second TIC readings over ~5 minutes.
These steps have been automated in Tac32Plus.

## Slide 15
From: Roberto Ambrosini, Tom Clark, Brian Corey, and Ed Himwich To: All IVS Stations Date: 1 May 2014
We recommend the following practices for management of the 1 PPS derived from the Maser and used as the station 1 PPS. Its synchronization with UTC as derived from the GPS 1 PPS offers a common timing reference for all VLBI stations worldwide. We refer to the difference in the epochs of the Maser and GPS 1 PPS signals, as measured by a counter, as the Maser/GPS offset, regardless of which signal occurs later.
Because it is evident that crossing zero time for the Maser/GPS offset should be carefully avoided (the counter would read the complement of one second of the desired delay, arithmetic processing of data by the counter not being recommended), we recommend keeping the offset at a small but significant distance from zero and its drift rate positive.
We also recommend keeping the time and frequency retuning of the Maser at a minimum, typically no more than once in a year.
This procedure offers: less work at the station, better modelling of the long term drift of the Maser, and a better chance to identify jumps in the offset. Here follow some practical recommendations for the Maser/GPS offset:
(1) Either the Maser 1 PPS or GPS 1 PPS can occur first.
(2) The offset should be significantly, at least a few microseconds, different from zero.
(3) The offset should not be too large, a useful upper limit might be on the order of 100 microseconds.
(4) The offset should be growing slowly, typically less than 0.1 microseconds/day.
(5) The offset should not be adjusted unnecessarily, no more often than once per year if possible.
(6) Items (2)-(5) are only recommendations and may not be feasible in some situations and do not need to replace existing successful practice at any station. However to the extent it is reasonable, stations should align themselves with these practices.
IVS Recommended Maser Timing Practices

## Slide 16
Recommendation (1) is a recognition that different stations have different preferences on which 1 PPS occurs first: Maser or GPS. Recommendations (2)-(4) are intended to minimize both the need to re-tune the Maser and the chances of the offset going through zero. Recommendation (5) is intended to make it easier to relate the offset data from one experiment to another.
For completeness, the following requirements (as opposed to recommendations) are listed for the FS log recorded offset between GPS and formatter 1PPS signals, the "GPS/FM offset". These requirements are necessary to allow correct interpretation of the offset data downstream. Please note that these requirements deal with the GPS/FM offset, which is related to, but different from Maser/GPS offset discussed above. In addition to the GPS/FM offset, stations can, and are encouraged to, record (appropriately labelled) additional available clock offset data, including the Maser/GPS offset, in their FS logs or separately. The requirements for the GPS/FM offset recorded in the FS logs:
(7) The offset is positive and small, i.e. close to (but not too close to) zero and NOT close to one second. If the recommendations (2)-(4) for the Maser/GPS offset above are used for that offset, they are likely to also be true for the GPS/Maser offset as well. In any event, the GPS/FM offset should not cross zero.
(8) The offset is recorded with either of two possible commands depending on how the counter is connected. The connections should be chosen to agree with (7) and: (A) If the counter is started by the GPS 1 PPS, use the "gps-fmout" command. This should be the case if the formatter output 1 PPS (usually determined by the Maser) is late. (B) If the counter is started by the fmout 1 PPS, use the "fmout-gps" command. This should be the case if the GPS 1 PPS is late. It will be necessary to change which command is used if which signal is late changes. This should not be needed if recommendations (2)-(4) for the Maser/GPS offset are followed.
(9) The offset counter does not use arithmetical processing. It just reports the "raw" difference in time between the start and stop signal. So for example, the small positive offset in (7) is not achieved by subtracting the raw difference from 1 second.
(10) The offset counter does not use averaging. This allows immediate detection of jumps. Averaging can be applied in post processing of the data.
(11) The offset must be measured at least once per scan in MIDOB. Additional measurements are acceptable as well.
IVS Recommended Maser Timing Practices

## Slide 17
Hydrogen Masers

## Slide 18
Credit: Microsemi MHM2010 Manual
Physics Package

## Slide 19
Credit: Microsemi MHM2010 Manual
Sigma Tau/NR Masers

## Slide 20
Sigma Tau MHM 2010
NR Maser
2 5MHz
2 10 MHz
2 1PPS
Maser Data
Sync Port
4 5MHz
2 1 PPS
Maser Data
Maser Outputs

## Slide 21
Timing Configuration

## Slide 22
Maser Data Monitoring
Frequency Data – TAC32 Plus
Data/Frequency Monitoring

## Slide 23
Hydrogen Pressure
Microprocessor batteries
Magnetics/Degaussing
Vacion pumps
Hydrogen gas
Frequency corrections
Troubleshooting/Routine Maintenance

## Slide 24
Power Outages
  Temperature instabilities-heater currents
  Loss of IF/VCO
  Backup Batteries
Microprocessor Failure
Power Supplies
Fuses
Troubleshooting/Routine Maintenance

## Slide 25: Now let’s discuss . . .
Week rollover may mean retiring old GPS receivers (Motorola VP, UT+, etc.)
  “We have legacy equipment using the Oncore VP. We have found that the VP receivers have a cutoff date after which the date reverts back 1024 weeks. The final VP sw ver 10 rev 0 has an end date of operation of 13 Aug 2016”
GPS receiver’s quantization error (“sawtooth”).
“Absolute” Receiver Calibration
New developments
  The SSR-M8T GNSS receiver
  Tac32Plus updates
  CNS Clock II functional improvements (NTP, Oscillator. PPS)

## Slide 26: An Example of 1PPS Quantization Error Correction
~26 nsec p-to-p
~1.5 nsec RMS noise (after applying quantization correction)

## Slide 27: What Causes the Quantization Error ?
For the older Oncore, F=9.54 MHz, so the 1/F quantization error has a range of +/- 52 nsec (104 nsec peak-to-peak)
The newer M12+ & M12M have F  40 MHz, so the quantization error has been reduced to +/- 13 nsec (26 nsec).

## Slide 28: VLBI’s Annoying Problem Caused by the Quantization Timing Error
When the formatter (Mark 5/6 sampler) needs to be reset, you have to feed it a 1PPS timing pulse to restart the internal VLBI clock. After it is started, it runs smoothly at a rate defined by the Maser’s 5/10 MHz.
The AVERAGE of the 1PPS pulses from the GPS receiver is “correct”, but any single pulse can be in error by ±13 nsec (or ±52 nsec with the older VP & UT Oncore receivers) because of the quantization error.
Once you have restarted the formatter with the noisy 1 PPS signal, you then measure the actual (GPS minus Formatter) time that you actually achieved.
Or, you can use the 1PPS from a CNS Clock II which has the quantization error removed.

## Slide 29: How can the quantization noise be eliminated ?
GPS
Timing
Receiver
Programmable Delay Line
with 250 psec steps
(Dallas/Maxim DS1023)
1PPS with
quantization noise
Serial Data
“Clean”
1PPS
RS-232
PIC Microprocessor generates the
correction for the NEXT 1PPS tick

## Slide 30: The Future is here now! The CNS Clock II
Data available on RS-232, USB 2.0, Ethernet LAN, RS-485 and solid state relay Ports.
Ethernet NTP Server for your LAN.
TNC GPS Antenna Connector.
Buffered 1 PPS outputs.
GPSDO 10 (or 5) MHz output.
High Performance PPS.
Steered OCXO.
Steered Oscillator Utility Functions.
Many Options: IRIG-B, etc.
1994 – 2004: the TAC
1PPS Sawtooth
Correction Option
Available Since January 2005, now at Revision J

## Slide 31: Does the hardware 1PPS correction work?

## Slide 32
Tac32Plus software simultaneously processes data from four Time Interval Counters and four CNS Clocks, writing 12 logs continuously.
Time Interval Counters compare the 1PPS from each CNS Clock (M12+) against the USNO’s
UTC time tick.
CNS Systems’ Test Bed at USNO Calibrating the “DC” Offset of M12+ receivers with 2.0 Firmware in 2002
We have observed that the ONCORE firmware evolution from 5.x  6.x  8.x  10.x has been accompanied by about 40 nsec of “DC” timing offsets.
Motorola tasked Rick to make the new M12+ receiver be correct.
This is the “Gold Standard” “A” receiver that we used for subsequent calibrations.

## Slide 33: Individual M12 Clock Performance “Gold” Receiver (A) average “DC” offset = -0.6 ns

## Slide 34: Comparing four M12+ Timing Receivers

## Slide 35: Trying to keep up with New Technology!
Around 2005, Motorola quit the GPS business and the M-12 design was licensed to iLotus in Singapore. The current iLotus timing receiver is called the M-12M.
Anticipating the need for a M-12 replacement, Rick & Art Sepin (Synergy) examined the Swiss-built uBlox LEA-6 receiver module. Because of the large installed base of Motorola/iLotus receivers, Rick developed a hybrid M-12 emulator; an M-12 sized board and a PIC μP to convert the uBlox binary command set to the Motorola @@ binary format.
More recently, this new receiver has been upgraded with improved hardware, firmware and the uBlox M8T GNSS module that supports multiple satellite systems, not just GPS. This is in the latest CNS Clock II product.

## Slide 36
An iLotus M-12M module.
The M12+ looks just the same
The Synergy SSR-M8T Receiver
Comparing an M12+, M-12M & uBlox LEA-6/M8
The uBlox LEA6T module

## Slide 37
“Gold”
Motorola M12+
iLotus
M12-M
Synergy LEA-6Ts
uBlox uBlox
Moto Native
Emul . Cmds
Maser 1PPS Distributor
4@ HP53132 Counters
The 4 Receiver test @ GGAO

## Slide 38: 10 days of 1 minute averages of Sigma-Tau 1pps tick to each of 4 rcvrs. Maser rate ~ 27.3 nsec/day Clock offsets ~ 8 μsec
Chart: 
Raw Measurements

## Slide 39
Chart: 
Removing 27.3 nsec/day H-Maser Rate

## Slide 40
Chart: 
Comparison of M-12s with "best" ublox Rcvr

## Slide 41
Modified Allan Deviation

## Slide 42
Small, low cost GPS receivers can provide timing needed for VLBI anywhere in the world. This is not a new statement, it’s been true since the 1990’s! See www.gpstime.com for “Timing for VLBI” notes from the IVS TOWs for more details.
The iLotus M12M we tested showed jumps at the 10 nsec level. Not sure if this is just a problem of this particular unit.
Existing designs based on Motorola/iLotus M12s should have no problem in making the change to uBlox by using the Synergy M12 emulator receivers.
The Synergy SSR receiver with either the uBlox LEA-6T (GPS only) or LEA-M8T (GNSS) is a superior product. In fact, the uBlox we tested were a factor ~5 BETTER than the old M12’s in all tests except for a “DC” bias ~30 nsec.
Conclusions

## Slide 43
Agilent has announced “End-of-Life” for the 53131 and 53132 counters that have been the standard VLBI Time Interval Counter.
The 53131/132 has been very easy to use since it has an RS232 printer port that streams the readings. TAC32 was built around this capability.
Agilent is recommending the 53230A as their suggested replacement for the 131/132.
Berkeley Nucleonics offers their model Model 1104 as an alternative.
Both these counters suggest their use as “Net Appliances” on the station LAN using their Ethernet ports.
More Obsolescence Problems

## Slide 44
CNS Clock II
HP/Agilent 53132A
Serial Port
Agilent 53230A
Ethernet
Berkeley Nucleonics Model 1104
Ethernet
Tac32Plus V2.7.11
Note: GPS time vs. HP5065A Rubidium CNS Systems’ time standard
Tac32Plus v2.7.16 Now Supports Time Interval Counters via Ethernet.

## Slide 45
TIC Setup is simple and familiar
53132A vs. BN1105
53132A vs. 53230A
Katie Pazamickas, Tom Clark & Rick Hamblly
@Haystack, May 2015
Tac32Plus v2.7.16 Now Supports Time Interval Counters via Ethernet.

## Slide 46
Support for Synergy SSR-M8T GNSS receiver
  GPS, BeiDou, GLONASS, QZSS. More coming soon.
  Label all satellite types in the signal strength bars.
Ability to upload new firmware into the Synergy SSR receivers.
Version number now in the main frame title bar.
Added the @@Wm "Change Model Number" command for the SSR receivers.
Many bug fixes and minor enhancements. More coming soon.
See www.cnssys.com/Tac32Plus/Tac32Plus.php.
Other Recent Tac32Plus Upgrades
Version 2.7.11 -> 2.7.16

## Slide 47
Connect to CNS Clock II via TCP/IP
Add GUI for satellite constellation selection on GNSS receivers like the SSR-M8T.
Investigate missing data in the BNC 1105 time interval counter log collected over a LAN. Reporting “Query INTERRUPTED” and “Query UNTERMINATED”.
Multi-Platform executables, especially Linux.
Other enhancements based on user feedback.
Contact Rick Hambly: rick@cnssys.com
Future Tac32Plus Enhancements: